HH1385

User Manual

HH1385

4-Channel Thermometer

User’s Manual

- Introduction

The HH1385 is a digital 4 input thermometer and data logger that works with any K, J, E, T, R, S, N type thermocouple temperature sensors. This unit offers higher accuracy and better uniformity across various thermocouple types. This device measures/stores over 16,800 temperature readings and allows up to 4 thermocouple connections that plug into a miniature female thermocouple receptacle at the base of the unit.

The user can easily set up the initial data logging parameters including thermocouple type, logging Intervals, high/low alarm settings, and desired temperature units (°C,°F or °K). The user may also download the stored data by plugging the meter straight into a PC’s USB port and running the easy-to-use Windows software.

# Accessories

1   User manual

6   1.5V AAA Carbon zinc battery

1   Carrying bag

1   9V AC to DC adaptor

- USB cable

# Safety Precaution

|  | Note! Please refer to this manual. Improper use may damage the meter and its components. |
| --- | --- |
|  | Complies with European Directive. |

- Do not operate in environments with flammable gas or humid environments

- Operating altitude: up to 2000M

- Operating environment: Indoor use; Pollution degree 2

- Clean with soft cloth when dirty, such as glasses cloth

- Do not clean with chemicals and other solvents

Class B – Equipment for use in all establishments other than domestic

Group 1 – RF energy generated is needed for the internal functioning

# Meter Description

- Alarm setup button / Auto record button

- Power button

- Backlight button

- K /J/ T / E/ R / S/ N type switching/manually record button.

- Hold button / Auto-shutdown button

- Interchange button for "T1 - T2" vs. "T1 & T2"

- Unit ° C / ° F / K switch button

- Maximum value, Minimum value, Average value, elapsed time / Setup button

- Relative value / offset value button

- Interchange button for "T3 - T4" vs. "T3 & T4"

- Time button / Read button

- Thermocouple input terminal T1/T2 T3/T4

- External power DC 9V input

- USB data output jack

- Battery cover and tilt stand

- LCD display

|  | - With DC-IN connected, appears on the screen. - Without heat-sensing wire connected, will appear on the screen. |
| --- | --- |

# Operation

## Power on or off

Press        button，turn on or turn off.

## Turn on backlight

Press       button to turn on or turn off the backlight.

＊The backlight will automatically turn off after being lit for 15 seconds.

## Alarm on or off

Press       to turn “on” or “off” the alarm, followed by "LIMIT" appears on the LCD.

This LCD shows the status "on" or "off" of the alarm function. It is possible to set the upper and lower limit of the temperature range by yourself. When the temperature reading exceeds the range, the buzzer will alarm until the temperature back to the range or the alarm function turned off.

(For the limit range setup, please enter Setup Function)

## Auto Recording

Press        and hold for ≥ 2 seconds to enable or disable the auto-recording function "LOG"

The LCD display will first show the number of logs for REC/MEM START mode and then

switch back to temperature mode after 1 second.

The auto-recording is such if the current number is 3, it will be increased to X by 1 at each

pressing. (the number of logs in HH1385 is up to 16800)

* Execute "LOG" to disable the auto-shutdown function. For recording for a long time,

please connect the meter to an external power supply.

## Thermocouple Type Selection (TYPE):

Press          to enable selection of different thermocouple types in a full cycle.

## Manual Record

Press       , followed by "REC" appears on the LCD, the log number of records will increase by +1 at each time, and one log will be recorded.

The manual records are up to 200 logs.

## Unit Switch for °C, °F and °K

Press        and hold for ≥ 2 seconds to switch the unit: Celsius (°C), Fahrenheit (°F), or Kelvin (°K).

## Data Hold

Press        to enable or disable the data hold function.

## Disable Auto Power Off Function

Press        and hold for 2 seconds to enable or disable the auto power off function.

## Relative Value Measurement

Press         to enter the relative:

Just use one thermocouple, the difference between the two temperatures can be measured.

For example, measure the first temperature as 25°C, and then press        to show 0°C on the LCD. Again, measure the second temperature. If the second temperature is measured as 30°C, the LCD will show the difference value 5°C (30-25=5°C).

Again, press        to disable the function.

The left shows the relative value, right shows the temperature value being measured.

Press the         /        and hold to switch the display between "T1 & T2" and "T3 & T4".

## Offset Value

The user can set the offset value to compensate for the error due to the working thermocouple wire.

Press         and hold for ≥ 2 seconds to set the offset value in OFFSET mode.

Press       or       to increase or decrease the offset value by the scale of 0.1°C, °F, or °K. The setting range is ±5°C, ±5K, or ±9°F.

*If to set offset value of another channels, press       and       to switch and select between "T1 & T2" and "T3 & T4" for adjustment.

Press          to save the setting value and exit the setting mode.

*If the value updated to be 0.1, "OFFSET" will continuously appear on the LCD. "OFFSET" will not disappear unless it is corrected back to 0.0.

## Change Display of T1-T2 vs. T1 & T2

Press         to enable or disable T1 – T2 .

The value shown on the left shows the difference of T1 – T2, and the one shown on the right is the T1 value which is measured presently.

Press and hold         for ≥ 2 seconds for the position change of T1 and T2.

## Change Display of T3-T4 vs. T3 & T4

Press          to show the value of T3 - T4.

Again, press         to disable the function.

The value shown on the left presents the difference of T3-T4, the one shown on the right is

the temperature measured through T3 channel.

Press and hold        for 2 seconds for the position change of T4 and T3 on the right.

## Time Display

Press     for < 1 second to switch the three date-time modes: Year/Month/Day →

Hour/Minute/Second → Cyclic displays each term by second.

## Read Record

Press       and hold for ≥ 2 seconds to read the records in "MEM" mode while LCD displays

MEM CALL.

Press       or        to read the previous or next record.

The three items of time, number of logs and temperature data are displayed on the LCD

and are auto-displayed cyclic by second.

It is used in the "MEM CALL" mode. Press        to exit this mode.

## Maximum /Minimum/ Average Value

Press         repeatedly to select maximum, minimum and average value of measured

data.

When MAX, MIN, and AVG are displayed at the same time, it represents the elapsed

time of the measurement, and the present temperature T1 / T3.

## Setup

Press           and hold for ≥ 2 seconds for SETUP

#### Step1. High/Low limit Value setting for alarm function (Limit)

Select Setup mode to the upper and lower limit of T1, T2, T3 and T4 (Setup Limit Hi、

Lo).

Press          to select and set the range (Hi & Lo) of T1, T2, T3, T4.

Press          to select the position to be modified. The position will be indicated to the

next one by each press.

Press       or       to set the plus-minus sign and temperature.

The setting of the upper and the lower limits are based on the TYPE range measured;

it may auto-identify whether it exceeds the limit. (The Setup function will fail if it exceeds

the TYPE range which is set by the user.)

Press           to save the settings and proceed to Step 2.

#### Step2.	Set Record Intervals

Press           or           to select the hour, minute, and second.

Press         or         to adjust the record interval.

Press           to save the settings and proceed to Step 3.

Step3.	Time Setting

Press         or         to select the year, month, day, hour, minute, and second.

Press        or         adjust the time.

Press and hold the button to enable speedy adjustment.

Press           to save the settings and proceed to Step 4.

#### Step4.	Clear the record in the memory

Press        to clear, "YES" will flash.

Press         to keep, "NO" will flash.

Press           to save the settings and complete all settings.

# Software Installation

- Supported operating systems: XP/Windows7/Windows 8.1/Windows10

# Install the desktop thermometer program (Please close virus scanning software when

# Executing.)

Place the CD into the CD/DVD-ROM player

Install the Temperature Meter T1 software

This will install Temperature Meter T1 Version 1.0.08 on your computer. It is recommended that you close all other applications before continuing. Click next to continue, or Cancel to exit Setup

## Finish Install and run Temperature Metert11.0.08 desktop

## software.

## (HH1385 is not connected to the PC)

Install Microsoft NET Framework Client Profile

Install Shield Wizard of PL2303

The Install Shield Wizard will install PL-2303 USB-to-Serial on your computer.

# Software manual

- Connect the USB cable included with a meter to the PC, as shown in the figure below.

Double click “Run Temperature Meter T1110” icon on the desktop

Select COM port connect to the PC

“USB”: PC is connected to a meter

“OL”: No Thermocouple connected

Year / Month / day   2019/11/27     Hour: Minute    01:07

Number of records / Total number of records    894/16800

Connection port COM5 (random change)

Select the “Start to record measured data Alarm function T1 / T2/T3/T4 Upper / lower limit setting (Alarm)

# Download Temperature Meter data from the meter’s memory

# Select the “Download recorded measurement record from / Temperature Meter”

## Save the record after the completion of the download

Save a new file (*.txt)

Top 5 highest data of T1 T2 T3 T4

Top 5 lowest data of T1 T2 T3 T4

Print T1 T2 T3 T4 temperature Chart

Change time and temperature range on the temperature chart

Change “record sampling time” (Min sample time 1 sec; Max sample time 86400sec)

Change “time to stop recording”

# General specifications

Display: 4-channel and 4-digit LCD

Unit: ℃ / ℉ / °K

Data hold (HOLD)

Auto-ranging

Backlight

Auto power (default 15 min) and disable auto power off

Maximum/minimum/mean value/measurement elapsed time

Alarm function

Overload display: “OL”

Input limit: Maximum input 24V DC or AC

Datalogging capacity 16,800 records

Save interval: 1 second~24 hour

Low battery detection

Battery: 1.5V×6 (LR03 SIZE AAA 1.5V)

Battery life: Approximately 100 hours

Operation temperature and humidity: 0°C to 50°C (32°F to 122°F), < 80%RH

Storage temperature and humidity:0°C to 50°C, relative humidity under 80%.

Weight: Approximately 330 grams

Dimensions: 168 (L) x 73 (W) x 35 (H) mm

The backlight will be continuously on if the meter is connected to the external power

- supply

AC to DC Adaptor

External AC 100~240V to DC 9V/0.5A adaptor. (Please pay attention to the polarity)

Voltage: DC9V (9.0 ~ 15.0 VDC MAX)

Current: ≧1000mA。

Plug: The pin in the center connects to the positive electrode and the outer case is a negative electrode

Diameter: 5.5mm; internal diameter: 2.1mm.

# Electrical specifications

# Maintenance or repair

When the When   “   ” symbol is displayed on the LCD, it means that there is insufficient power; please change the battery immediately in order to ensure its accuracy.

Do not place the meter in locations that have a high temperature, humidity or that are exposed to direct sunlight.

Remember to turn off the power after usage; remove the battery if not used for a long period of time in order to prevent battery leakage and causing damages to internal components.

When the instrument failure, only by the authorized service provider or return the original repair.

# Battery replacement

- Turn off the power.

- Open the frame and battery cover at the back of the meter, remove the batteries.

- Please insert a new AAA battery according to the polarities.

- Put the battery cover and frame back in place.

# End of life disposal

Note: This symbol indicates that the meter and its accessories must be

separated and processed properly.